EMG-Analyzer-Pro

Project Documentation  |  Automated quality validation of multi-channel surface EMG recordings

# 1. Overview

What it is. A Python notebook (Analysis.ipynb, runs in Google Colab or Jupyter) that takes a raw surface-EMG recording and returns one PDF report with a clear PASS / CAUTION / FAIL verdict, a 0–100 score for each channel, and a prioritised list of fixes.

The problem. sEMG signals are microvolt-level and easily ruined by loose electrodes, mains hum, amplifier clipping or dropped packets. Bad recordings silently wreck gesture-recognition models, and the damage is usually found only after hours of data collection and training.

The solution. Check the recording immediately after capture, with objective rules, so a bad session is re-recorded the same day.

Who it is for: researchers and students collecting gesture data with an 8-channel, 500 Hz device (prosthetics, rehabilitation, gesture control). No coding is needed – you only press Run and upload two files.

Key design decisions and why

| Decision | Reason |
| --- | --- |
| Rule-based scoring, not machine learning | Every deduction can be explained and traced to a threshold. Users can tune thresholds, and no labelled “good/bad” training set is required. |
| One notebook, one PDF output | Zero installation in Colab, a single file to archive or share with a supervisor, and every figure sits beside its explanation. |
| Control Panel with sliders, no code edits | Lets non-programmers adapt the tool to another device, mains frequency (50/60 Hz) or gesture set. |
| AI is optional and fenced in | Offline rule-based text is the default. Gemini or Claude only write wording; numbers, tables, plots and verdicts always come from code (see Section 6). |

# 2. How It Works

The recording passes through five stages. Each stage answers one question, and a failure at an early stage explains problems later on.

| # | Stage | What is measured | Why it matters |
| --- | --- | --- | --- |
| 1 | Timing and packets | Sampling interval, backward/duplicate timestamps, missing packet numbers, coverage vs. expected samples | Gaps and jitter distort filters and spectra; short sessions have too few repetitions. |
| 2 | Channel health | Mean, std, peak-to-peak, % at ADC max (clipping), % at zero, flat 1-s windows, channel correlation | Finds dead, disconnected or saturated electrodes in seconds. |
| 3 | Signal quality | Gesture/rest RMS ratio, SNR, mains-band power, spectrogram | Shows whether muscle activity actually rises above the noise floor. |
| 4 | Data readiness | Repeatability (CV) across trials, label balance, window purity, Fisher ratio, Mahalanobis distance | Predicts whether a classifier can separate the gestures. |
| 5 | Verdict and report | Scorecard, 9 recording-level checks, recommended actions, PDF | Converts numbers into one decision plus actions. |

# 3. Quick Start

- Open Analysis.ipynb in Google Colab (no installation) or Jupyter. Missing libraries install automatically.

- Run the Control Panel cell. The defaults suit the 8-channel, 500 Hz device. Change only what differs (for example mains 60 Hz or the rest label).

- Run all remaining cells (Runtime → Run after). When asked, upload emg_data_*.csv and, optionally, metadata_*.txt.

- Download EMG_Signal_Analysis_Report.pdf, which starts with the verdict and the action list.

Input files

| File | Contents |
| --- | --- |
| emg_data_*.csv (required) | Timestamp (ms), optional Packet_Number, EMG1–EMG8 as raw ADC counts, Label (gesture; “rest” is the baseline), Trial (repetition index). Column names are auto-matched, so naming differences are tolerated. |
| metadata_*.txt (optional) | Sampling rate, expected sample count. Without it, the rate is estimated from timestamps and coverage checks are skipped. |

Key Control Panel settings

| Setting | Default | Reason for the default |
| --- | --- | --- |
| Mains frequency | 50 Hz | Set 60 Hz in the Americas; the wrong value hides real hum. |
| ADC resolution | 12 bit | Needed to detect clipping at the true maximum count. |
| Window / stride | 200 / 50 ms | Windows under about 300 ms keep real-time control responsive; 75% overlap gives plenty of samples. |
| Min std / peak-to-peak | 10 / 50 counts | Below this the channel is effectively flat. |
| Max saturation / zeros | 0.1% / 1.0% | Even brief clipping removes signal peaks, so the tolerance is tight. |
| Min gesture/rest ratio | 1.2 | Muscle must be visibly more active than rest. |
| Mains warning | 15% of power | Hum above this masks the 40–70 Hz motor-unit range. |

# 4. Methods and Reasoning

Each method was chosen because it is standard in the EMG literature, cheap to compute and easy to interpret.

| Method | How it works | Why this choice |
| --- | --- | --- |
| Band-pass 20–200 Hz | 4th-order Butterworth, applied forward and backward (zero-phase) | Most EMG energy lies in 20–200 Hz. Below 20 Hz is motion artefact; above 200 Hz is noise. Zero-phase keeps gesture onsets correctly aligned in time. |
| Linear envelope | Rectify, then 5 Hz low-pass | Gives a smooth picture of contraction strength, so a human can verify activity visually. |
| SNR (dB) | 20·log10(active RMS / rest RMS) | Rest is the real noise floor of this setup, so no separate noise recording is needed. |
| Welch PSD and mains % | Averaged spectrum; power within ±2 Hz of 50/60 Hz divided by total power | Averaging gives a stable estimate for short segments. A power ratio is independent of signal amplitude. |
| Median / mean frequency | Frequency splitting power in half / power-weighted centroid | A downward drift across trials is the classic sign of muscle fatigue, and it explains accuracy that falls over time. |
| Fast window RMS | Cumulative sum of squares, so each window costs O(1) | Hundreds of thousands of windows compute almost instantly in a notebook. |
| Label purity | Keep only windows with a single label | Windows crossing Rest→Fist mix two states and would blur the class statistics. |
| Fisher ratio | Between-class / within-class variance per channel | Ranks channels by usefulness, so weak electrodes can be repositioned or dropped. |
| Mahalanobis distance | Distance between gesture centroids scaled by pooled covariance | Accounts for correlated channels (cross-talk), which Euclidean distance ignores. Guide: below 1 heavy overlap, 1–2 moderate, 2–3 separable, above 3 well separated. |
| Coefficient of variation | Std / mean of RMS across trials | Reveals inconsistent effort or electrode shift. About 20% or less is good; above 50% is suspect. |

# 5. Scoring and Verdict

Channel score. Each channel starts at 100 and loses points per problem. GOOD is 80 or more, FAIR is 50–79, POOR is below 50 (consider excluding the channel).

| Problem | Trigger | Points | Suggested fix |
| --- | --- | --- | --- |
| LOW_SIGNAL_VARIATION | std < 10 counts | 30 | Check electrode and wiring. |
| LOW_DYNAMIC_RANGE | peak-to-peak < 50 | 20 | Check gain and skin contact. |
| ADC_SATURATION | > 0.1% at max (> 5% costs 45) | 30 / 45 | Lower gain; check cable. |
| ZERO_VALUES | > 1% at zero (> 10% costs 45) | 25 / 45 | Re-seat electrode; check ground. |
| WEAK_REST_GESTURE_SEPARATION | gesture/rest RMS < 1.2 | 20 | Move electrode onto the muscle belly. |
| HIGH_MAINS_INTERFERENCE | > 15% power near mains | 10 | Skin prep, shielding, notch filter. |
| FLATLINE_SEGMENTS | > 5% of 1-s windows flat | 20 | Check for intermittent disconnection. |

Recording verdict. The overall result is the worst outcome among the recording-level checks:

| Verdict | Triggered by |
| --- | --- |
| FAIL | Fewer than half of the channels score 50 or more; sample coverage under 50%; fewer than two non-rest gestures. |
| CAUTION | Any channel below 80; coverage under 95%; more than 1% packets missing; no rest label; only one trial; a gesture pair with distance below 1; smallest gesture under half the size of the largest. |
| PASS | None of the above. |

Why these rules. Penalties are weighted by how much each fault harms later analysis: dead or clipped channels (30–45) outweigh mains noise (10), which can be repaired by filtering. The thresholds are engineering heuristics and are tunable in the Control Panel.

# 6. Report and AI Safeguards

The PDF contains the verdict, scorecard, actions, every table and figure, and a plain-language explanation next to each result. Explanation text comes from one of three providers:

| Provider | Cost / needs | Best used when |
| --- | --- | --- |
| Rule-based (default) | Free, offline, no key | Sensitive data, or when a reproducible result is needed. Fixed templates filled with computed numbers, so nothing can be invented. |
| Gemini (2.5 Flash) | Free tier, API key in Colab Secrets | Richer wording at no cost. |
| Claude | Paid, API key in Colab Secrets | Highest-quality narrative. |

Anti-hallucination protection (Gemini and Claude): the model sees only numbers computed by the notebook; every number in its answer is checked against that data; an answer with unverifiable numbers is requested again once; any sentence still containing one is deleted; tables, plots, scores and the verdict are never written by AI. Privacy: Gemini and Claude receive your computed results (and figures if enabled), so use Rule-based for confidential data.

# 7. Limitations and Future Work

- Heuristic thresholds. They are tuned for one device and are not clinically validated; recalibrate for new hardware.

- Separability is not classifier accuracy. Distances predict data quality, not model performance.

- Runs on one recording at a time and depends on Colab upload/download.

- Next steps: batch processing of many sessions, a notch-filter option, comparison across sessions to track fatigue, and a standalone command-line version.

Key references: De Luca (1997), J. Applied Biomechanics 13(2); Merletti & Parker (2004), Electromyography, Wiley/IEEE; Welch (1967), IEEE Trans. Audio Electroacoust. 15(2); Hermens et al. (2000), SENIAM, J. Electromyogr. Kinesiol. 10(5).